## Asset_Cal_Info
Sensor Code | startDateTime | stopDateTime | Sensor.uid | Sensor Serial Number | Calibration Cofficient Name | Calibration Cofficient Value | Notes | Source | Undefined Sensors
PHSENA | 2013-02-07 |  | ATAPL-58337-00001 | P0059 | CC_ea434 | 17533 | mCP molar absorptivities provided by vendor, specific to a reagent bag with a defined shelflife. | From RS03AXPS
PHSENA | 2013-02-07 |  | ATAPL-58337-00001 | P0059 | CC_eb434 | 2229 | mCP molar absorptivities provided by vendor, specific to a reagent bag with a defined shelflife.
PHSENA | 2013-02-07 |  | ATAPL-58337-00001 | P0059 | CC_ea578 | 101 | mCP molar absorptivities provided by vendor, specific to a reagent bag with a defined shelflife.
PHSENA | 2013-02-07 |  | ATAPL-58337-00001 | P0059 | CC_eb578 | 38502 | mCP molar absorptivities provided by vendor, specific to a reagent bag with a defined shelflife.
PHSENA | 2013-02-07 |  | ATAPL-58337-00001 | P0059 | CC_ind_slp | 1 | indicator impurity offset correction factor [unitless]
PHSENA | 2013-02-07 |  | ATAPL-58337-00001 | P0059 | CC_ind_off | 0 | indicator impurity slope correction factor [unitless]
PHSENA | 2013-02-07 |  | ATAPL-58337-00001 | P0059 | CC_psal | 35 | Constant value defined in: def ph_calc_phwater(ref, light, therm, ea434, eb434, ea578, eb578, ind_slp, ind_off, psal=35.0)
PHSEND | 2015-01-01 |  | A00096 | P0068 | CC_ea434 | 17533 | mCP molar absorptivities provided by vendor, specific to a reagent bag with a defined shelflife. | From CP01CNSM-00003
PHSEND | 2015-01-01 |  | A00096 | P0068 | CC_ea578 | 101 | mCP molar absorptivities provided by vendor, specific to a reagent bag with a defined shelflife.
PHSEND | 2015-01-01 |  | A00096 | P0068 | CC_eb434 | 2229 | mCP molar absorptivities provided by vendor, specific to a reagent bag with a defined shelflife.
PHSEND | 2015-01-01 |  | A00096 | P0068 | CC_eb578 | 38502 | mCP molar absorptivities provided by vendor, specific to a reagent bag with a defined shelflife.
PHSEND | 2015-01-01 |  | A00096 | P0068 | CC_ind_off | 0 | indicator impurity offset correction factor [unitless]
PHSEND | 2015-01-01 |  | A00096 | P0068 | CC_ind_slp | 1 | indicator impurity slope correction factor [unitless]
PHSEND | 2015-01-01 |  | A00096 | P0068 | CC_psal | 35 | Constant value defined in: def ph_calc_phwater(ref, light, therm, ea434, eb434, ea578, eb578, ind_slp, ind_off, psal=35.0)
PHSEND | 2015-01-01 |  | A00097 | P0069 | CC_ea434 | 17533 | mCP molar absorptivities provided by vendor, specific to a reagent bag with a defined shelflife. | From CP01CNSM-00003
PHSEND | 2015-01-01 |  | A00097 | P0069 | CC_ea578 | 101 | mCP molar absorptivities provided by vendor, specific to a reagent bag with a defined shelflife.
PHSEND | 2015-01-01 |  | A00097 | P0069 | CC_eb434 | 2229 | mCP molar absorptivities provided by vendor, specific to a reagent bag with a defined shelflife.
PHSEND | 2015-01-01 |  | A00097 | P0069 | CC_eb578 | 38502 | mCP molar absorptivities provided by vendor, specific to a reagent bag with a defined shelflife.
PHSEND | 2015-01-01 |  | A00097 | P0069 | CC_ind_off | 0 | indicator impurity offset correction factor [unitless]
PHSEND | 2015-01-01 |  | A00097 | P0069 | CC_ind_slp | 1 | indicator impurity slope correction factor [unitless]
PHSEND | 2015-01-01 |  | A00097 | P0069 | CC_psal | 35 | Constant value defined in: def ph_calc_phwater(ref, light, therm, ea434, eb434, ea578, eb578, ind_slp, ind_off, psal=35.0)
PHSEND | 2015-01-01 |  | A00119 | P0070 | CC_ea434 | 17533 |  | From GA01SUMO-00001
PHSEND | 2015-01-01 |  | A00119 | P0070 | CC_eb434 | 2229
PHSEND | 2015-01-01 |  | A00119 | P0070 | CC_ea578 | 101
PHSEND | 2015-01-01 |  | A00119 | P0070 | CC_eb578 | 38502
PHSEND | 2015-01-01 |  | A00119 | P0070 | CC_ind_slp | 1
PHSEND | 2015-01-01 |  | A00119 | P0070 | CC_ind_off | 0
PHSEND | 2015-01-01 |  | A00119 | P0070 | CC_psal | 35 | Constant
PHSEND | 2014-06-01 |  | A00118 | P0071 | CC_ea434 | 17533 | From vendor cal sheet | From GI01SUMO-00001
PHSEND | 2014-06-01 |  | A00118 | P0071 | CC_eb434 | 2229 | From vendor cal sheet
PHSEND | 2014-06-01 |  | A00118 | P0071 | CC_ea578 | 101 | From vendor cal sheet
PHSEND | 2014-06-01 |  | A00118 | P0071 | CC_eb578 | 38502 | From vendor cal sheet
PHSEND | 2014-06-01 |  | A00118 | P0071 | CC_ind_slp | 1 | Apply to all CGSN and RSN pH sensors; provided by the vendor as part of their Matlab code; DPS author wanted set as input to the code.
PHSEND | 2014-06-01 |  | A00118 | P0071 | CC_ind_off | 0 | Apply to all CGSN and RSN pH sensors; provided by the vendor as part of their Matlab code; DPS author wanted set as input to the code.
PHSEND | 2014-06-01 |  | A00118 | P0071 | CC_psal | 35 | Default per <ph_calc_phwater(ref, light, therm, ea434, eb434, ea578, eb578, ind_slp, ind_off, psal=35.0)>
PHSEND | 2016-01-01 |  | A00118 | P0071 | CC_ea434 | 17372 | From vendor cal sheet | From GI01SUMO-00003
PHSEND | 2016-01-01 |  | A00118 | P0071 | CC_eb434 | 2284 | From vendor cal sheet
PHSEND | 2016-01-01 |  | A00118 | P0071 | CC_ea578 | 94.1 | From vendor cal sheet
PHSEND | 2016-01-01 |  | A00118 | P0071 | CC_eb578 | 38676 | From vendor cal sheet
PHSEND | 2016-01-01 |  | A00118 | P0071 | CC_ind_slp | 1 | Apply to all CGSN and RSN pH sensors; provided by the vendor as part of their Matlab code; DPS author wanted set as input to the code.
PHSEND | 2016-01-01 |  | A00118 | P0071 | CC_ind_off | 0 | Apply to all CGSN and RSN pH sensors; provided by the vendor as part of their Matlab code; DPS author wanted set as input to the code.
PHSEND | 2016-01-01 |  | A00118 | P0071 | CC_psal | 35 | Default per <ph_calc_phwater(ref, light, therm, ea434, eb434, ea578, eb578, ind_slp, ind_off, psal=35.0)>
PHSEND | 2013-01-29 |  | A00249 | P0072 | CC_ea434 | 17533 |  | From CE01ISSM-00001, CE02SHSM-00002
PHSEND | 2013-01-29 |  | A00249 | P0072 | CC_ea578 | 101
PHSEND | 2013-01-29 |  | A00249 | P0072 | CC_eb434 | 2229
PHSEND | 2013-01-29 |  | A00249 | P0072 | CC_eb578 | 38502
PHSEND | 2013-01-29 |  | A00249 | P0072 | CC_ind_off | 0
PHSEND | 2013-01-29 |  | A00249 | P0072 | CC_ind_slp | 1
PHSEND | 2013-01-29 |  | A00249 | P0072 | CC_psal | 35 | Default per <ph_calc_phwater(ref, light, therm, ea434, eb434, ea578, eb578, ind_slp, ind_off, psal=35.0)>
PHSENF | 2014-06-01 |  | A00160 | P0079 | CC_ea434 | 17533 | From vendor cal sheet | From GI03FLMA-00001
PHSENF | 2014-06-01 |  | A00160 | P0079 | CC_eb434 | 2229 | From vendor cal sheet
PHSENF | 2014-06-01 |  | A00160 | P0079 | CC_ea578 | 101 | From vendor cal sheet
PHSENF | 2014-06-01 |  | A00160 | P0079 | CC_eb578 | 38502 | From vendor cal sheet
PHSENF | 2014-06-01 |  | A00160 | P0079 | CC_ind_slp | 1 | Apply to all CGSN and RSN pH sensors; provided by the vendor as part of their Matlab code; DPS author wanted set as input to the code.
PHSENF | 2014-06-01 |  | A00160 | P0079 | CC_ind_off | 0 | Apply to all CGSN and RSN pH sensors; provided by the vendor as part of their Matlab code; DPS author wanted set as input to the code.
PHSENF | 2014-06-01 |  | A00160 | P0079 | CC_psal | 35 | Default per <ph_calc_phwater(ref, light, therm, ea434, eb434, ea578, eb578, ind_slp, ind_off, psal=35.0)>
PHSENF | 2016-01-01 |  | A00160 | P0079 | CC_ea434 | 17372 |  | From GP03FLMA-00004
PHSENF | 2016-01-01 |  | A00160 | P0079 | CC_eb434 | 2284.1
PHSENF | 2016-01-01 |  | A00160 | P0079 | CC_ea578 | 94.1
PHSENF | 2016-01-01 |  | A00160 | P0079 | CC_eb578 | 38676.5
PHSENF | 2016-01-01 |  | A00160 | P0079 | CC_ind_slp | 1
PHSENF | 2016-01-01 |  | A00160 | P0079 | CC_ind_off | 0
PHSENF | 2016-01-01 |  | A00160 | P0079 | CC_psal | 35
PHSENF | 2013-01-01 |  | A00161 | P0080 | CC_ea434 | 17533 |  | From GP03FLMA-00001
PHSENF | 2013-01-01 |  | A00161 | P0080 | CC_eb434 | 2229
PHSENF | 2013-01-01 |  | A00161 | P0080 | CC_ea578 | 101
PHSENF | 2013-01-01 |  | A00161 | P0080 | CC_eb578 | 38502
PHSENF | 2013-01-01 |  | A00161 | P0080 | CC_ind_slp | 1
PHSENF | 2013-01-01 |  | A00161 | P0080 | CC_ind_off | 0
PHSENF | 2013-01-01 |  | A00161 | P0080 | CC_psal | 35 | Default per <ph_calc_phwater(ref, light, therm, ea434, eb434, ea578, eb578, ind_slp, ind_off, psal=35.0)>
PHSENF | 2015-01-01 |  | A00161 | P0080 | CC_ea434 | 17533 |  | From GP03FLMA-00003
PHSENF | 2015-01-01 |  | A00161 | P0080 | CC_eb434 | 2229
PHSENF | 2015-01-01 |  | A00161 | P0080 | CC_ea578 | 101
PHSENF | 2015-01-01 |  | A00161 | P0080 | CC_eb578 | 38502
PHSENF | 2015-01-01 |  | A00161 | P0080 | CC_ind_slp | 1
PHSENF | 2015-01-01 |  | A00161 | P0080 | CC_ind_off | 0
PHSENF | 2015-01-01 |  | A00161 | P0080 | CC_psal | 35 | Constant
PHSENF | 2013-01-01 |  | A00162 | P0081 | CC_ea434 | 17533 |  | From GP03FLMB-00001
PHSENF | 2013-01-01 |  | A00162 | P0081 | CC_eb434 | 2229
PHSENF | 2013-01-01 |  | A00162 | P0081 | CC_ea578 | 101
PHSENF | 2013-01-01 |  | A00162 | P0081 | CC_eb578 | 38502
PHSENF | 2013-01-01 |  | A00162 | P0081 | CC_ind_slp | 1
PHSENF | 2013-01-01 |  | A00162 | P0081 | CC_ind_off | 0
PHSENF | 2013-01-01 |  | A00162 | P0081 | CC_psal | 35 | Default per <ph_calc_phwater(ref, light, therm, ea434, eb434, ea578, eb578, ind_slp, ind_off, psal=35.0)>
PHSEND | 2013-03-08 |  | A00261 | P0083 | CC_ea434 | 17533 |  | From CE01ISSM-00001, CE01ISSM-00003
PHSEND | 2013-03-08 |  | A00261 | P0083 | CC_ea578 | 101
PHSEND | 2013-03-08 |  | A00261 | P0083 | CC_eb434 | 2229
PHSEND | 2013-03-08 |  | A00261 | P0083 | CC_eb578 | 38502
PHSEND | 2013-03-08 |  | A00261 | P0083 | CC_ind_off | 0
PHSEND | 2013-03-08 |  | A00261 | P0083 | CC_ind_slp | 1
PHSEND | 2013-03-08 |  | A00261 | P0083 | CC_psal | 35 | Default per <ph_calc_phwater(ref, light, therm, ea434, eb434, ea578, eb578, ind_slp, ind_off, psal=35.0)>
PHSEND | 2016-01-01 |  | A00261 | P0083 | CC_ea434 | 17372 |  | From CE09OSSM-00003
PHSEND | 2016-01-01 |  | A00261 | P0083 | CC_ea578 | 94.1
PHSEND | 2016-01-01 |  | A00261 | P0083 | CC_eb434 | 2284
PHSEND | 2016-01-01 |  | A00261 | P0083 | CC_eb578 | 38676
PHSEND | 2016-01-01 |  | A00261 | P0083 | CC_ind_off | 0
PHSEND | 2016-01-01 |  | A00261 | P0083 | CC_ind_slp | 1
PHSEND | 2016-01-01 |  | A00261 | P0083 | CC_psal | 35 | Default per <ph_calc_phwater(ref, light, therm, ea434, eb434, ea578, eb578, ind_slp, ind_off, psal=35.0)>
PHSEND | 2012-11-05 |  | A00250 | P0084 | CC_ea434 | 17533 |  | From CE01ISSM-00001, CE06ISSM-00002
PHSEND | 2012-11-05 |  | A00250 | P0084 | CC_ea578 | 101
PHSEND | 2012-11-05 |  | A00250 | P0084 | CC_eb434 | 2229
PHSEND | 2012-11-05 |  | A00250 | P0084 | CC_eb578 | 38502
PHSEND | 2012-11-05 |  | A00250 | P0084 | CC_ind_off | 0
PHSEND | 2012-11-05 |  | A00250 | P0084 | CC_ind_slp | 1
PHSEND | 2012-11-05 |  | A00250 | P0084 | CC_psal | 35 | Default per <ph_calc_phwater(ref, light, therm, ea434, eb434, ea578, eb578, ind_slp, ind_off, psal=35.0)>
PHSEND | 2016-01-01 |  | A00250 | P0084 | CC_ea434 | 17372 |  | From CE01ISSM-00004
PHSEND | 2016-01-01 |  | A00250 | P0084 | CC_ea578 | 94.1
PHSEND | 2016-01-01 |  | A00250 | P0084 | CC_eb434 | 2284
PHSEND | 2016-01-01 |  | A00250 | P0084 | CC_eb578 | 38676
PHSEND | 2016-01-01 |  | A00250 | P0084 | CC_ind_off | 0
PHSEND | 2016-01-01 |  | A00250 | P0084 | CC_ind_slp | 1
PHSEND | 2016-01-01 |  | A00250 | P0084 | CC_psal | 35 | Default per <ph_calc_phwater(ref, light, therm, ea434, eb434, ea578, eb578, ind_slp, ind_off, psal=35.0)>
PHSEND | 2013-03-12 |  | A00262 | P0085 | CC_ea434 | 17533 |  | From CE06ISSM-00001, CE07SHSM-00002
PHSEND | 2013-03-12 |  | A00262 | P0085 | CC_ea578 | 101
PHSEND | 2013-03-12 |  | A00262 | P0085 | CC_eb434 | 2229
PHSEND | 2013-03-12 |  | A00262 | P0085 | CC_eb578 | 38502
PHSEND | 2013-03-12 |  | A00262 | P0085 | CC_ind_off | 0
PHSEND | 2013-03-12 |  | A00262 | P0085 | CC_ind_slp | 1
PHSEND | 2013-03-12 |  | A00262 | P0085 | CC_psal | 35 | Default per <ph_calc_phwater(ref, light, therm, ea434, eb434, ea578, eb578, ind_slp, ind_off, psal=35.0)>
PHSEND | 2013-03-14 |  | A00266 | P0086 | CC_ea434 | 17533 |  | From CE01ISSM-00003
PHSEND | 2013-03-14 |  | A00266 | P0086 | CC_ea578 | 101
PHSEND | 2013-03-14 |  | A00266 | P0086 | CC_eb434 | 2229
PHSEND | 2013-03-14 |  | A00266 | P0086 | CC_eb578 | 38502
PHSEND | 2013-03-14 |  | A00266 | P0086 | CC_ind_off | 0
PHSEND | 2013-03-14 |  | A00266 | P0086 | CC_ind_slp | 1
PHSEND | 2013-03-14 |  | A00266 | P0086 | CC_psal | 35 | Default per <ph_calc_phwater(ref, light, therm, ea434, eb434, ea578, eb578, ind_slp, ind_off, psal=35.0)>
PHSEND | 2016-01-01 |  | A00266 | P0086 | CC_ea434 | 17372 |  | From CE01ISSM-00003
PHSEND | 2016-01-01 |  | A00266 | P0086 | CC_ea578 | 94.1
PHSEND | 2016-01-01 |  | A00266 | P0086 | CC_eb434 | 2284
PHSEND | 2016-01-01 |  | A00266 | P0086 | CC_eb578 | 38676
PHSEND | 2016-01-01 |  | A00266 | P0086 | CC_ind_off | 0
PHSEND | 2016-01-01 |  | A00266 | P0086 | CC_ind_slp | 1
PHSEND | 2016-01-01 |  | A00266 | P0086 | CC_psal | 35 | Default per <ph_calc_phwater(ref, light, therm, ea434, eb434, ea578, eb578, ind_slp, ind_off, psal=35.0)>
[9 empty] | PHSEND | 2013-03-12 | A00264 | P0087 | CC_ea434
[9 empty] | PHSEND | 2013-03-12 | A00264 | P0087 | CC_eb434
[9 empty] | PHSEND | 2013-03-12 | A00264 | P0087 | CC_ea578
[9 empty] | PHSEND | 2013-03-12 | A00264 | P0087 | CC_eb578
[9 empty] | PHSEND | 2013-03-12 | A00264 | P0087 | CC_ind_slp
[9 empty] | PHSEND | 2013-03-12 | A00264 | P0087 | CC_ind_off
[9 empty] | PHSEND | 2013-03-12 | A00264 | P0087 | CC_psal | Default per <ph_calc_phwater(ref, light, therm, ea434, eb434, ea578, eb578, ind_slp, ind_off, psal=35.0)>
PHSEND | 2013-03-12 |  | A00263 | P0088 | CC_ea434 | 17533 |  | From CE07SHSM-00001
PHSEND | 2013-03-12 |  | A00263 | P0088 | CC_ea578 | 101
PHSEND | 2013-03-12 |  | A00263 | P0088 | CC_eb434 | 2229
PHSEND | 2013-03-12 |  | A00263 | P0088 | CC_eb578 | 38502
PHSEND | 2013-03-12 |  | A00263 | P0088 | CC_ind_off | 0
PHSEND | 2013-03-12 |  | A00263 | P0088 | CC_ind_slp | 1
PHSEND | 2013-03-12 |  | A00263 | P0088 | CC_psal | 35 | Constant
PHSEND | 2016-01-01 |  | A00263 | P0088 | CC_ea434 | 17372 |  | From CE01ISSM-00005
PHSEND | 2016-01-01 |  | A00263 | P0088 | CC_ea578 | 94.1
PHSEND | 2016-01-01 |  | A00263 | P0088 | CC_eb434 | 2284
PHSEND | 2016-01-01 |  | A00263 | P0088 | CC_eb578 | 38676
PHSEND | 2016-01-01 |  | A00263 | P0088 | CC_ind_off | 0
PHSEND | 2016-01-01 |  | A00263 | P0088 | CC_ind_slp | 1
PHSEND | 2016-01-01 |  | A00263 | P0088 | CC_psal | 35 | Constant
PHSEND | 2013-06-01 |  | A00388 | P0095 | CC_ea434 | 17533 |  | From CP01CNSM-00001
PHSEND | 2013-06-01 |  | A00388 | P0095 | CC_ea578 | 101
PHSEND | 2013-06-01 |  | A00388 | P0095 | CC_eb434 | 2229
PHSEND | 2013-06-01 |  | A00388 | P0095 | CC_eb578 | 38502
PHSEND | 2013-06-01 |  | A00388 | P0095 | CC_ind_off | 0
PHSEND | 2013-06-01 |  | A00388 | P0095 | CC_ind_slp | 1 | Constant
PHSEND | 2013-06-01 |  | A00388 | P0095 | CC_psal | 35
PHSEND | 2015-01-01 |  | A00388 | P0095 | CC_ea434 | 17533 |  | From CP03ISSM-00002
PHSEND | 2015-01-01 |  | A00388 | P0095 | CC_ea578 | 101
PHSEND | 2015-01-01 |  | A00388 | P0095 | CC_eb434 | 2229
PHSEND | 2015-01-01 |  | A00388 | P0095 | CC_eb578 | 38502
PHSEND | 2015-01-01 |  | A00388 | P0095 | CC_ind_off | 0
PHSEND | 2015-01-01 |  | A00388 | P0095 | CC_ind_slp | 1
PHSEND | 2015-01-01 |  | A00388 | P0095 | CC_psal | 35 | Constant
PHSEND | 2016-01-01 |  | A00388 | P0095 | CC_ea434 | 17372 |  | From CP03ISSM-00004
PHSEND | 2016-01-01 |  | A00388 | P0095 | CC_ea578 | 94.1
PHSEND | 2016-01-01 |  | A00388 | P0095 | CC_eb434 | 2284
PHSEND | 2016-01-01 |  | A00388 | P0095 | CC_eb578 | 38676
PHSEND | 2016-01-01 |  | A00388 | P0095 | CC_ind_off | 1
PHSEND | 2016-01-01 |  | A00388 | P0095 | CC_ind_slp | 0
PHSEND | 2016-01-01 |  | A00388 | P0095 | CC_psal | 35 | Constant
PHSEND | 2013-06-01 |  | A00389 | P0096 | CC_ea434 | 17533 |  | From CP01CNSM-00001
PHSEND | 2013-06-01 |  | A00389 | P0096 | CC_ea578 | 101
PHSEND | 2013-06-01 |  | A00389 | P0096 | CC_eb434 | 2229
PHSEND | 2013-06-01 |  | A00389 | P0096 | CC_eb578 | 38502
PHSEND | 2013-06-01 |  | A00389 | P0096 | CC_ind_off | 0
PHSEND | 2013-06-01 |  | A00389 | P0096 | CC_ind_slp | 1
PHSEND | 2013-06-01 |  | A00389 | P0096 | CC_psal | 35 | Constant
PHSEND | 2016-01-01 |  | A00389 | P0096 | CC_ea434 | 17533 |  | From CP04OSSM-00004
PHSEND | 2016-01-01 |  | A00389 | P0096 | CC_ea578 | 101
PHSEND | 2016-01-01 |  | A00389 | P0096 | CC_eb434 | 2229
PHSEND | 2016-01-01 |  | A00389 | P0096 | CC_eb578 | 38502
PHSEND | 2016-01-01 |  | A00389 | P0096 | CC_ind_off | 1
PHSEND | 2016-01-01 |  | A00389 | P0096 | CC_ind_slp | 0
PHSEND | 2016-01-01 |  | A00389 | P0096 | CC_psal | 35 | Constant
PHSEND | 2014-06-01 |  | A00422 | P0098 | CC_ea434 | 17533 |  | From CP04OSSM-00001
PHSEND | 2014-06-01 |  | A00422 | P0098 | CC_ea578 | 101
PHSEND | 2014-06-01 |  | A00422 | P0098 | CC_eb434 | 2229
PHSEND | 2014-06-01 |  | A00422 | P0098 | CC_eb578 | 38502
PHSEND | 2014-06-01 |  | A00422 | P0098 | CC_ind_off | 0
PHSEND | 2014-06-01 |  | A00422 | P0098 | CC_ind_slp | 1
PHSEND | 2014-06-01 |  | A00422 | P0098 | CC_psal | 35 | Constant
PHSEND | 2016-01-01 |  | A00422 | P0098 | CC_ea434 | 17533 |  | From CP04OSSM-00004
PHSEND | 2016-01-01 |  | A00422 | P0098 | CC_ea578 | 101
PHSEND | 2016-01-01 |  | A00422 | P0098 | CC_eb434 | 2229
PHSEND | 2016-01-01 |  | A00422 | P0098 | CC_eb578 | 38502
PHSEND | 2016-01-01 |  | A00422 | P0098 | CC_ind_off | 1
PHSEND | 2016-01-01 |  | A00422 | P0098 | CC_ind_slp | 0
PHSEND | 2016-01-01 |  | A00422 | P0098 | CC_psal | 35 | Constant
PHSEND | 2014-06-01 |  | A00423 | P0099 | CC_ea434 | 17533 |  | From CP04OSSM-00001
PHSEND | 2014-06-01 |  | A00423 | P0099 | CC_ea578 | 101
PHSEND | 2014-06-01 |  | A00423 | P0099 | CC_eb434 | 2229
PHSEND | 2014-06-01 |  | A00423 | P0099 | CC_eb578 | 38502
PHSEND | 2014-06-01 |  | A00423 | P0099 | CC_ind_off | 0
PHSEND | 2014-06-01 |  | A00423 | P0099 | CC_ind_slp | 1
PHSEND | 2014-06-01 |  | A00423 | P0099 | CC_psal | 35 | Constant
PHSEND | 2016-06-01 |  | A00423 | P0099 | CC_ea434 | 17533 |  | From CP03ISSM-00004
PHSEND | 2016-06-01 |  | A00423 | P0099 | CC_ea578 | 101
PHSEND | 2016-06-01 |  | A00423 | P0099 | CC_eb434 | 2229
PHSEND | 2016-06-01 |  | A00423 | P0099 | CC_eb578 | 38502
PHSEND | 2016-06-01 |  | A00423 | P0099 | CC_ind_off | 0
PHSEND | 2016-06-01 |  | A00423 | P0099 | CC_ind_slp | 1
PHSEND | 2016-06-01 |  | A00423 | P0099 | CC_psal | 35 | Constant
PHSEND | 2015-01-01 |  | A00392 | P0100 | CC_ea434 | 17533 |  | From CP03ISSM-00002
PHSEND | 2015-01-01 |  | A00392 | P0100 | CC_ea578 | 101
PHSEND | 2015-01-01 |  | A00392 | P0100 | CC_eb434 | 2229
PHSEND | 2015-01-01 |  | A00392 | P0100 | CC_eb578 | 38502
PHSEND | 2015-01-01 |  | A00392 | P0100 | CC_ind_off | 0
PHSEND | 2015-01-01 |  | A00392 | P0100 | CC_ind_slp | 1
PHSEND | 2015-01-01 |  | A00392 | P0100 | CC_psal | 35 | Constant
PHSEND | 2016-04-01 |  | A00392 | P0100 | CC_ea434 | 17533 |  | From CP01CNSM-00005
PHSEND | 2016-04-01 |  | A00392 | P0100 | CC_ea578 | 101
PHSEND | 2016-04-01 |  | A00392 | P0100 | CC_eb434 | 2229
PHSEND | 2016-04-01 |  | A00392 | P0100 | CC_eb578 | 38502
PHSEND | 2016-04-01 |  | A00392 | P0100 | CC_ind_off | 1
PHSEND | 2016-04-01 |  | A00392 | P0100 | CC_ind_slp | 0
PHSEND | 2016-04-01 |  | A00392 | P0100 | CC_psal | 35 | Constant
PHSEND | 2014-06-01 |  | A00567 | P0101 | CC_ea434 | 17533 | From vendor cal sheet | From GI01SUMO-00001
PHSEND | 2014-06-01 |  | A00567 | P0101 | CC_eb434 | 2229 | From vendor cal sheet
PHSEND | 2014-06-01 |  | A00567 | P0101 | CC_ea578 | 101 | From vendor cal sheet
PHSEND | 2014-06-01 |  | A00567 | P0101 | CC_eb578 | 38502 | From vendor cal sheet
PHSEND | 2014-06-01 |  | A00567 | P0101 | CC_ind_slp | 1 | Apply to all CGSN and RSN pH sensors; provided by the vendor as part of their Matlab code; DPS author wanted set as input to the code.
PHSEND | 2014-06-01 |  | A00567 | P0101 | CC_ind_off | 0 | Apply to all CGSN and RSN pH sensors; provided by the vendor as part of their Matlab code; DPS author wanted set as input to the code.
PHSEND | 2014-06-01 |  | A00567 | P0101 | CC_psal | 35 | Default per <ph_calc_phwater(ref, light, therm, ea434, eb434, ea578, eb578, ind_slp, ind_off, psal=35.0)>
PHSEND | 2016-01-01 |  | A00567 | P0101 | CC_ea434 | 17372 | From vendor cal sheet | From GI01SUMO-00003
PHSEND | 2016-01-01 |  | A00567 | P0101 | CC_eb434 | 2284 | From vendor cal sheet
PHSEND | 2016-01-01 |  | A00567 | P0101 | CC_ea578 | 94.1 | From vendor cal sheet
PHSEND | 2016-01-01 |  | A00567 | P0101 | CC_eb578 | 38676 | From vendor cal sheet
PHSEND | 2016-01-01 |  | A00567 | P0101 | CC_ind_slp | 1 | Apply to all CGSN and RSN pH sensors; provided by the vendor as part of their Matlab code; DPS author wanted set as input to the code.
PHSEND | 2016-01-01 |  | A00567 | P0101 | CC_ind_off | 0 | Apply to all CGSN and RSN pH sensors; provided by the vendor as part of their Matlab code; DPS author wanted set as input to the code.
PHSEND | 2016-01-01 |  | A00567 | P0101 | CC_psal | 35 | Default per <ph_calc_phwater(ref, light, therm, ea434, eb434, ea578, eb578, ind_slp, ind_off, psal=35.0)>
PHSENF | 2014-01-01 |  | A01501 | P0102 | CC_ea434 | 17533 |  | From GP03FLMB-00002
PHSENF | 2014-01-01 |  | A01501 | P0102 | CC_eb434 | 2229
PHSENF | 2014-01-01 |  | A01501 | P0102 | CC_ea578 | 101
PHSENF | 2014-01-01 |  | A01501 | P0102 | CC_eb578 | 38502
PHSENF | 2014-01-01 |  | A01501 | P0102 | CC_ind_slp | 1
PHSENF | 2014-01-01 |  | A01501 | P0102 | CC_ind_off | 0
PHSENF | 2014-01-01 |  | A01501 | P0102 | CC_psal | 35 | Default per <ph_calc_phwater(ref, light, therm, ea434, eb434, ea578, eb578, ind_slp, ind_off, psal=35.0)>
PHSENF | 2014-01-01 |  | A01502 | P0103 | CC_ea434 | 17533 |  | From GP03FLMA-00002
PHSENF | 2014-01-01 |  | A01502 | P0103 | CC_eb434 | 2229
PHSENF | 2014-01-01 |  | A01502 | P0103 | CC_ea578 | 101
PHSENF | 2014-01-01 |  | A01502 | P0103 | CC_eb578 | 38502
PHSENF | 2014-01-01 |  | A01502 | P0103 | CC_ind_slp | 1
PHSENF | 2014-01-01 |  | A01502 | P0103 | CC_ind_off | 0
PHSENF | 2014-01-01 |  | A01502 | P0103 | CC_psal | 35 | Default per <ph_calc_phwater(ref, light, therm, ea434, eb434, ea578, eb578, ind_slp, ind_off, psal=35.0)>
PHSENF | 2016-01-01 |  | A01502 | P0103 | CC_ea434 | 17533 |  | From GI03FLMA-00003
PHSENF | 2016-01-01 |  | A01502 | P0103 | CC_eb434 | 2229
PHSENF | 2016-01-01 |  | A01502 | P0103 | CC_ea578 | 101
PHSENF | 2016-01-01 |  | A01502 | P0103 | CC_eb578 | 38502
PHSENF | 2016-01-01 |  | A01502 | P0103 | CC_ind_slp | 1
PHSENF | 2016-01-01 |  | A01502 | P0103 | CC_ind_off | 0
PHSENF | 2016-01-01 |  | A01502 | P0103 | CC_psal | 35 | Default per <ph_calc_phwater(ref, light, therm, ea434, eb434, ea578, eb578, ind_slp, ind_off, psal=35.0)>
PHSENF | 2015-01-01 |  | A00468 | P0104 | CC_ea434 | 17533 | From vendor cal sheet | From GI03FLMB-00001
PHSENF | 2015-01-01 |  | A00468 | P0104 | CC_eb434 | 2229 | From vendor cal sheet
PHSENF | 2015-01-01 |  | A00468 | P0104 | CC_ea578 | 101 | From vendor cal sheet
PHSENF | 2015-01-01 |  | A00468 | P0104 | CC_eb578 | 38502 | From vendor cal sheet
PHSENF | 2015-01-01 |  | A00468 | P0104 | CC_ind_slp | 1 | Apply to all CGSN and RSN pH sensors; provided by the vendor as part of their Matlab code; DPS author wanted set as input to the code.
PHSENF | 2015-01-01 |  | A00468 | P0104 | CC_ind_off | 0 | Apply to all CGSN and RSN pH sensors; provided by the vendor as part of their Matlab code; DPS author wanted set as input to the code.
PHSENF | 2015-01-01 |  | A00468 | P0104 | CC_psal | 35 | Default per <ph_calc_phwater(ref, light, therm, ea434, eb434, ea578, eb578, ind_slp, ind_off, psal=35.0)>
PHSENF | 2016-01-01 |  | A00468 | P0104 | CC_ea434 | 17372 | From vendor cal sheet | From GI03FLMB-00003
PHSENF | 2016-01-01 |  | A00468 | P0104 | CC_eb434 | 2284 | From vendor cal sheet
PHSENF | 2016-01-01 |  | A00468 | P0104 | CC_ea578 | 94.1 | From vendor cal sheet
PHSENF | 2016-01-01 |  | A00468 | P0104 | CC_eb578 | 38676 | From vendor cal sheet
PHSENF | 2016-01-01 |  | A00468 | P0104 | CC_ind_slp | 1 | Apply to all CGSN and RSN pH sensors; provided by the vendor as part of their Matlab code; DPS author wanted set as input to the code.
PHSENF | 2016-01-01 |  | A00468 | P0104 | CC_ind_off | 0 | Apply to all CGSN and RSN pH sensors; provided by the vendor as part of their Matlab code; DPS author wanted set as input to the code.
PHSENF | 2016-01-01 |  | A00468 | P0104 | CC_psal | 35 | Default per <ph_calc_phwater(ref, light, therm, ea434, eb434, ea578, eb578, ind_slp, ind_off, psal=35.0)>
PHSENF | 2014-06-01 |  | A00490 | P0105 | CC_ea434 | 17533 | From vendor cal sheet | From GP03FLMB-00003
PHSENF | 2014-06-01 |  | A00490 | P0105 | CC_eb434 | 2229 | From vendor cal sheet
PHSENF | 2014-06-01 |  | A00490 | P0105 | CC_ea578 | 101 | From vendor cal sheet
PHSENF | 2014-06-01 |  | A00490 | P0105 | CC_eb578 | 38502 | From vendor cal sheet
PHSENF | 2014-06-01 |  | A00490 | P0105 | CC_ind_slp | 1 | Apply to all CGSN and RSN pH sensors; provided by the vendor as part of their Matlab code; DPS author wanted set as input to the code.
PHSENF | 2014-06-01 |  | A00490 | P0105 | CC_ind_off | 0 | Apply to all CGSN and RSN pH sensors; provided by the vendor as part of their Matlab code; DPS author wanted set as input to the code.
PHSENF | 2014-06-01 |  | A00490 | P0105 | CC_psal | 35 | Default per <ph_calc_phwater(ref, light, therm, ea434, eb434, ea578, eb578, ind_slp, ind_off, psal=35.0)>
PHSENA | 2013-02-07 |  | ATAPL-58337-00004 | P0110 | CC_ea434 | 17533 | From vendor cal sheet | From RS03AXPS
PHSENA | 2013-02-07 |  | ATAPL-58337-00004 | P0110 | CC_eb434 | 2229 | From vendor cal sheet
PHSENA | 2013-02-07 |  | ATAPL-58337-00004 | P0110 | CC_ea578 | 101 | From vendor cal sheet
PHSENA | 2013-02-07 |  | ATAPL-58337-00004 | P0110 | CC_eb578 | 38502 | From vendor cal sheet
PHSENA | 2013-02-07 |  | ATAPL-58337-00004 | P0110 | CC_ind_slp | 0.9698 | Apply to all CGSN and RSN pH sensors; provided by the vendor as part of their Matlab code; DPS author wanted set as input to the code.
PHSENA | 2013-02-07 |  | ATAPL-58337-00004 | P0110 | CC_ind_off | 0.2484 | Apply to all CGSN and RSN pH sensors; provided by the vendor as part of their Matlab code; DPS author wanted set as input to the code.
PHSENA | 2013-02-07 |  | ATAPL-58337-00004 | P0110 | CC_psal | 35 | Default per <ph_calc_phwater(ref, light, therm, ea434, eb434, ea578, eb578, ind_slp, ind_off, psal=35.0)>
PHSENA | 2013-02-07 |  | ATAPL-58337-00003 | P0111 | CC_ea434 | 17533 |  | From RS01SBPS
PHSENA | 2013-02-07 |  | ATAPL-58337-00003 | P0111 | CC_eb434 | 2229
PHSENA | 2013-02-07 |  | ATAPL-58337-00003 | P0111 | CC_ea578 | 101
PHSENA | 2013-02-07 |  | ATAPL-58337-00003 | P0111 | CC_eb578 | 38502
PHSENA | 2013-02-07 |  | ATAPL-58337-00003 | P0111 | CC_ind_slp | 1
PHSENA | 2013-02-07 |  | ATAPL-58337-00003 | P0111 | CC_ind_off | 0
PHSENA | 2013-02-07 |  | ATAPL-58337-00003 | P0111 | CC_psal | 35
PHSENA | 2013-02-07 |  | ATAPL-58337-00002 | P0112 | CC_ea434 | 17533 |  | From RS03AXPS
PHSENA | 2013-02-07 |  | ATAPL-58337-00002 | P0112 | CC_eb434 | 2229
PHSENA | 2013-02-07 |  | ATAPL-58337-00002 | P0112 | CC_ea578 | 101
PHSENA | 2013-02-07 |  | ATAPL-58337-00002 | P0112 | CC_eb578 | 38502
PHSENA | 2013-02-07 |  | ATAPL-58337-00002 | P0112 | CC_ind_slp | 0.9698
PHSENA | 2013-02-07 |  | ATAPL-58337-00002 | P0112 | CC_ind_off | 0.2484
PHSENA | 2013-02-07 |  | ATAPL-58337-00002 | P0112 | CC_psal | 35
PHSENA | 2013-02-07 |  | ATAPL-58337-00005 | P0113 | CC_ea434 | 17533 | From vendor cal sheet | From RS01SBPS
PHSENA | 2013-02-07 |  | ATAPL-58337-00005 | P0113 | CC_eb434 | 2229 | From vendor cal sheet
PHSENA | 2013-02-07 |  | ATAPL-58337-00005 | P0113 | CC_ea578 | 101 | From vendor cal sheet
PHSENA | 2013-02-07 |  | ATAPL-58337-00005 | P0113 | CC_eb578 | 38502 | From vendor cal sheet
PHSENA | 2013-02-07 |  | ATAPL-58337-00005 | P0113 | CC_ind_slp | 1 | Apply to all CGSN and RSN pH sensors; provided by the vendor as part of their Matlab code; DPS author wanted set as input to the code.
PHSENA | 2013-02-07 |  | ATAPL-58337-00005 | P0113 | CC_ind_off | 0 | Apply to all CGSN and RSN pH sensors; provided by the vendor as part of their Matlab code; DPS author wanted set as input to the code.
PHSENA | 2013-02-07 |  | ATAPL-58337-00005 | P0113 | CC_psal | 35 | Default per <ph_calc_phwater(ref, light, therm, ea434, eb434, ea578, eb578, ind_slp, ind_off, psal=35.0)>
PHSEND | 2014-01-10 |  | A00629 | P0115 | CC_ea434 | 17533 |  | From CE04OSPS-00001
PHSEND | 2014-01-10 |  | A00629 | P0115 | CC_eb434 | 2229
PHSEND | 2014-01-10 |  | A00629 | P0115 | CC_ea578 | 101
PHSEND | 2014-01-10 |  | A00629 | P0115 | CC_eb578 | 38502
PHSEND | 2014-01-10 |  | A00629 | P0115 | CC_ind_slp | 1
PHSEND | 2014-01-10 |  | A00629 | P0115 | CC_ind_off | 0
PHSEND | 2014-01-10 |  | A00629 | P0115 | CC_psal | 35
PHSEND | 2014-01-23 |  | A00627 | P0116 | CC_ea434 | 17533 |  | From CE02SHBP-00001
PHSEND | 2014-01-23 |  | A00627 | P0116 | CC_eb434 | 2229
PHSEND | 2014-01-23 |  | A00627 | P0116 | CC_ea578 | 101
PHSEND | 2014-01-23 |  | A00627 | P0116 | CC_eb578 | 38502
PHSEND | 2014-01-23 |  | A00627 | P0116 | CC_ind_slp | 1
PHSEND | 2014-01-23 |  | A00627 | P0116 | CC_ind_off | 0
PHSEND | 2014-01-23 |  | A00627 | P0116 | CC_psal | 35
PHSEND | 2014-04-29 |  | A00648 | P0117 | CC_ea434 | 17533 |  | From CE04OSBP-00001
PHSEND | 2014-04-29 |  | A00648 | P0117 | CC_eb434 | 2229
PHSEND | 2014-04-29 |  | A00648 | P0117 | CC_ea578 | 101
PHSEND | 2014-04-29 |  | A00648 | P0117 | CC_eb578 | 38502
PHSEND | 2014-04-29 |  | A00648 | P0117 | CC_ind_slp | 1
PHSEND | 2014-04-29 |  | A00648 | P0117 | CC_ind_off | 0
PHSEND | 2014-04-29 |  | A00648 | P0117 | CC_psal | 35
PHSEND | 2014-01-27 |  | A00661 | P0118 | CC_ea434 | 17533 |  | From CE02SHSM-00001
PHSEND | 2014-01-27 |  | A00661 | P0118 | CC_ea578 | 101
PHSEND | 2014-01-27 |  | A00661 | P0118 | CC_eb434 | 2229
PHSEND | 2014-01-27 |  | A00661 | P0118 | CC_eb578 | 38502
PHSEND | 2014-01-27 |  | A00661 | P0118 | CC_ind_off | 0
PHSEND | 2014-01-27 |  | A00661 | P0118 | CC_ind_slp | 1
PHSEND | 2014-01-27 |  | A00661 | P0118 | CC_psal | 35 | Constant
PHSEND | 2014-01-30 |  | A00797 | P0119 | CC_ea434 | 17533 |  | From CE09OSSM-00001
PHSEND | 2014-01-30 |  | A00797 | P0119 | CC_ea578 | 101
PHSEND | 2014-01-30 |  | A00797 | P0119 | CC_eb434 | 2229
PHSEND | 2014-01-30 |  | A00797 | P0119 | CC_eb578 | 38502
PHSEND | 2014-01-30 |  | A00797 | P0119 | CC_ind_off | 0
PHSEND | 2014-01-30 |  | A00797 | P0119 | CC_ind_slp | 1
PHSEND | 2014-01-30 |  | A00797 | P0119 | CC_psal | 35 | Constant
PHSEND | 2014-01-30 |  | A00797 | P0119 | CC_ea434 | 17372 |  | From CE06ISSM-00004
PHSEND | 2014-01-30 |  | A00797 | P0119 | CC_ea578 | 94.1
PHSEND | 2014-01-30 |  | A00797 | P0119 | CC_eb434 | 2284
PHSEND | 2014-01-30 |  | A00797 | P0119 | CC_eb578 | 38676
PHSEND | 2014-01-30 |  | A00797 | P0119 | CC_ind_off | 0
PHSEND | 2014-01-30 |  | A00797 | P0119 | CC_ind_slp | 1
PHSEND | 2014-01-30 |  | A00797 | P0119 | CC_psal | 35 | Constant
PHSEND | 2014-04-29 |  | A00649 | P0120 | CC_ea434 | 17533 |  | From CE07SHSM-00001
PHSEND | 2014-04-29 |  | A00649 | P0120 | CC_ea578 | 101
PHSEND | 2014-04-29 |  | A00649 | P0120 | CC_eb434 | 2229
PHSEND | 2014-04-29 |  | A00649 | P0120 | CC_eb578 | 38502
PHSEND | 2014-04-29 |  | A00649 | P0120 | CC_ind_off | 0
PHSEND | 2014-04-29 |  | A00649 | P0120 | CC_ind_slp | 1
PHSEND | 2014-04-29 |  | A00649 | P0120 | CC_psal | 35 | Constant
PHSEND | 2014-01-21 |  | A00650 | P0121 | CC_ea434 | 17533 |  | From CE01ISSM-00002, CE09OSSM-00002
PHSEND | 2014-01-21 |  | A00650 | P0121 | CC_eb434 | 101
PHSEND | 2014-01-21 |  | A00650 | P0121 | CC_ea578 | 2229
PHSEND | 2014-01-21 |  | A00650 | P0121 | CC_eb578 | 38502
PHSEND | 2014-01-21 |  | A00650 | P0121 | CC_ind_slp | 0
PHSEND | 2014-01-21 |  | A00650 | P0121 | CC_ind_off | 1
PHSEND | 2014-01-21 |  | A00650 | P0121 | CC_psal | 35 | Constant
PHSEND | 2014-01-27 |  | A00660 | P0122 | CC_ea434 | 17533 |  | From CE09OSSM-00001
PHSEND | 2014-01-27 |  | A00660 | P0122 | CC_ea578 | 101
PHSEND | 2014-01-27 |  | A00660 | P0122 | CC_eb434 | 2229
PHSEND | 2014-01-27 |  | A00660 | P0122 | CC_eb578 | 38502
PHSEND | 2014-01-27 |  | A00660 | P0122 | CC_ind_off | 0
PHSEND | 2014-01-27 |  | A00660 | P0122 | CC_ind_slp | 1
PHSEND | 2014-01-27 |  | A00660 | P0122 | CC_psal | 35 | Constant
PHSEND | 2014-02-24 |  | A01042 | P0123 | CC_ea434 | 17533 |  | From CE01ISSM-00002, CE07SHSM-00002
PHSEND | 2014-02-24 |  | A01042 | P0123 | CC_ea578 | 101
PHSEND | 2014-02-24 |  | A01042 | P0123 | CC_eb434 | 2229
PHSEND | 2014-02-24 |  | A01042 | P0123 | CC_eb578 | 38502
PHSEND | 2014-02-24 |  | A01042 | P0123 | CC_ind_off | 0
PHSEND | 2014-02-24 |  | A01042 | P0123 | CC_ind_slp | 1 | Constant
PHSEND | 2014-02-24 |  | A01042 | P0123 | CC_psal | 35
PHSEND | 2014-01-27 |  | A00662 | P0124 | CC_ea434 | 17533 |  | From CE04OSSM-00001
PHSEND | 2014-01-27 |  | A00662 | P0124 | CC_ea578 | 101
PHSEND | 2014-01-27 |  | A00662 | P0124 | CC_eb434 | 2229
PHSEND | 2014-01-27 |  | A00662 | P0124 | CC_eb578 | 38502
PHSEND | 2014-01-27 |  | A00662 | P0124 | CC_ind_off | 0
PHSEND | 2014-01-27 |  | A00662 | P0124 | CC_ind_slp | 1
PHSEND | 2014-01-27 |  | A00662 | P0124 | CC_psal | 35 | Constant
PHSEND | 2014-02-12 |  | A00799 | P0125 | CC_ea434 | 17533 |  | From CE06ISSM-00002
PHSEND | 2014-02-12 |  | A00799 | P0125 | CC_ea578 | 101
PHSEND | 2014-02-12 |  | A00799 | P0125 | CC_eb434 | 2229
PHSEND | 2014-02-12 |  | A00799 | P0125 | CC_eb578 | 38502
PHSEND | 2014-02-12 |  | A00799 | P0125 | CC_ind_off | 0
PHSEND | 2014-02-12 |  | A00799 | P0125 | CC_ind_slp | 1
PHSEND | 2014-02-12 |  | A00799 | P0125 | CC_psal | 35
PHSEND | 2016-01-01 |  | A00799 | P0125 | CC_ea434 | 17372 |  | From CE01ISSM-00005
PHSEND | 2016-01-01 |  | A00799 | P0125 | CC_ea578 | 94.1
PHSEND | 2016-01-01 |  | A00799 | P0125 | CC_eb434 | 2284
PHSEND | 2016-01-01 |  | A00799 | P0125 | CC_eb578 | 38676
PHSEND | 2016-01-01 |  | A00799 | P0125 | CC_ind_off | 0
PHSEND | 2016-01-01 |  | A00799 | P0125 | CC_ind_slp | 1
PHSEND | 2016-01-01 |  | A00799 | P0125 | CC_psal | 35
PHSENA | 2014-09-19 |  | N00684 | P0131 | CC_ea434 | 17533 |  | From CE04OSPS-00001
PHSENA | 2014-09-19 |  | N00684 | P0131 | CC_eb434 | 2229
PHSENA | 2014-09-19 |  | N00684 | P0131 | CC_ea578 | 101
PHSENA | 2014-09-19 |  | N00684 | P0131 | CC_eb578 | 38502
PHSENA | 2014-09-19 |  | N00684 | P0131 | CC_ind_slp | 1
PHSENA | 2014-09-19 |  | N00684 | P0131 | CC_ind_off | 0
PHSENA | 2014-09-19 |  | N00684 | P0131 | CC_psal | 35 | Constant
PHSEND | 2014-09-19 |  | A01046 | P0132 | CC_ea434 | 17533 |  | From CE04OSBP-00002
PHSEND | 2014-09-19 |  | A01046 | P0132 | CC_eb434 | 2229
PHSEND | 2014-09-19 |  | A01046 | P0132 | CC_ea578 | 101
PHSEND | 2014-09-19 |  | A01046 | P0132 | CC_eb578 | 38502
PHSEND | 2014-09-19 |  | A01046 | P0132 | CC_ind_slp | 1
PHSEND | 2014-09-19 |  | A01046 | P0132 | CC_ind_off | 0
PHSEND | 2014-09-19 |  | A01046 | P0132 | CC_psal | 35 | Constant
PHSENA | 2013-02-07 |  | ATAPL-58337-00006 | P0133 | CC_ea434 | 17533 |  | From RS01SBPS
PHSENA | 2013-02-07 |  | ATAPL-58337-00006 | P0133 | CC_eb434 | 2229
PHSENA | 2013-02-07 |  | ATAPL-58337-00006 | P0133 | CC_ea578 | 101
PHSENA | 2013-02-07 |  | ATAPL-58337-00006 | P0133 | CC_eb578 | 38502
PHSENA | 2013-02-07 |  | ATAPL-58337-00006 | P0133 | CC_ind_slp | 1
PHSENA | 2013-02-07 |  | ATAPL-58337-00006 | P0133 | CC_ind_off | 0
PHSENA | 2013-02-07 |  | ATAPL-58337-00006 | P0133 | CC_psal | 35 | Constant
PHSENE | 2014-09-01 |  | A00897 | P0137 | CC_ea434 | 17533 |  | From GS01SUMO-00001
PHSENE | 2014-09-01 |  | A00897 | P0137 | CC_eb434 | 2229
PHSENE | 2014-09-01 |  | A00897 | P0137 | CC_ea578 | 101
PHSENE | 2014-09-01 |  | A00897 | P0137 | CC_eb578 | 38502
PHSENE | 2014-09-01 |  | A00897 | P0137 | CC_ind_slp | 1
PHSENE | 2014-09-01 |  | A00897 | P0137 | CC_ind_off | 0
PHSENE | 2014-09-01 |  | A00897 | P0137 | CC_psal | 35 | Constant
PHSEND | 2015-06-01 |  | A00916 | P0138 | CC_ea434 | 17533 |  | From GI01SUMO-00002
PHSEND | 2015-06-01 |  | A00916 | P0138 | CC_eb434 | 2229
PHSEND | 2015-06-01 |  | A00916 | P0138 | CC_ea578 | 101
PHSEND | 2015-06-01 |  | A00916 | P0138 | CC_eb578 | 38502
PHSEND | 2015-06-01 |  | A00916 | P0138 | CC_ind_slp | 1
PHSEND | 2015-06-01 |  | A00916 | P0138 | CC_ind_off | 0
PHSEND | 2015-06-01 |  | A00916 | P0138 | CC_psal | 35 | Constant
PHSEND | 2015-06-01 |  | A00898 | P0139 | CC_ea434 | 17533 |  | From GI01SUMO-00002
PHSEND | 2015-06-01 |  | A00898 | P0139 | CC_eb434 | 2229
PHSEND | 2015-06-01 |  | A00898 | P0139 | CC_ea578 | 101
PHSEND | 2015-06-01 |  | A00898 | P0139 | CC_eb578 | 38502
PHSEND | 2015-06-01 |  | A00898 | P0139 | CC_ind_slp | 1
PHSEND | 2015-06-01 |  | A00898 | P0139 | CC_ind_off | 0
PHSEND | 2015-06-01 |  | A00898 | P0139 | CC_psal | 35 | Constant
PHSEND | 2015-01-01 |  | A00917 | P0140 | CC_ea434 | 17533 |  | From GA01SUMO-00001
PHSEND | 2015-01-01 |  | A00917 | P0140 | CC_eb434 | 2229
PHSEND | 2015-01-01 |  | A00917 | P0140 | CC_ea578 | 101
PHSEND | 2015-01-01 |  | A00917 | P0140 | CC_eb578 | 38502
PHSEND | 2015-01-01 |  | A00917 | P0140 | CC_ind_slp | 1
PHSEND | 2015-01-01 |  | A00917 | P0140 | CC_ind_off | 0
PHSEND | 2015-01-01 |  | A00917 | P0140 | CC_psal | 35 | Constant
PHSEND | 2015-01-01 |  | N00501 | P0141 | CC_ea434 | 17533 |  | From GA03FLMA-00001
PHSEND | 2015-01-01 |  | N00501 | P0141 | CC_eb434 | 2229
PHSEND | 2015-01-01 |  | N00501 | P0141 | CC_ea578 | 101
PHSEND | 2015-01-01 |  | N00501 | P0141 | CC_eb578 | 38502
PHSEND | 2015-01-01 |  | N00501 | P0141 | CC_ind_slp | 1
PHSEND | 2015-01-01 |  | N00501 | P0141 | CC_ind_off | 0
PHSEND | 2015-01-01 |  | N00501 | P0141 | CC_psal | 35 | Constant
PHSENF | 2014-09-01 |  | A00972 | P0143 | CC_ea434 | 17533 |  | From GP03FLMB-00004
PHSENF | 2014-09-01 |  | A00972 | P0143 | CC_eb434 | 2229
PHSENF | 2014-09-01 |  | A00972 | P0143 | CC_ea578 | 101
PHSENF | 2014-09-01 |  | A00972 | P0143 | CC_eb578 | 38502
PHSENF | 2014-09-01 |  | A00972 | P0143 | CC_ind_slp | 1
PHSENF | 2014-09-01 |  | A00972 | P0143 | CC_ind_off | 0
PHSENF | 2014-09-01 |  | A00972 | P0143 | CC_psal | 35 | Constant
PHSENF | 2014-09-01 |  | A00973 | P0144 | CC_ea434 | 17533 |  | From GS03FLMB-00001
PHSENF | 2014-09-01 |  | A00973 | P0144 | CC_eb434 | 2229
PHSENF | 2014-09-01 |  | A00973 | P0144 | CC_ea578 | 101
PHSENF | 2014-09-01 |  | A00973 | P0144 | CC_eb578 | 38502
PHSENF | 2014-09-01 |  | A00973 | P0144 | CC_ind_slp | 1
PHSENF | 2014-09-01 |  | A00973 | P0144 | CC_ind_off | 0
PHSENF | 2014-09-01 |  | A00973 | P0144 | CC_psal | 35 | Constant
PHSENF | 2014-09-01 |  | A01500 | P0147 | CC_ea434 | 17533 |  | From GS03FLMA-00001
PHSENF | 2014-09-01 |  | A01500 | P0147 | CC_eb434 | 2229
PHSENF | 2014-09-01 |  | A01500 | P0147 | CC_ea578 | 101
PHSENF | 2014-09-01 |  | A01500 | P0147 | CC_eb578 | 38502
PHSENF | 2014-09-01 |  | A01500 | P0147 | CC_ind_slp | 1
PHSENF | 2014-09-01 |  | A01500 | P0147 | CC_ind_off | 0
PHSENF | 2014-09-01 |  | A01500 | P0147 | CC_psal | 35 | Constant
PHSEND | 2013-04-25 |  | ATOSU-58337-00011 | P0153 | CC_ea434 | 17533 |  | From CE04OSPS-00002
PHSEND | 2013-04-25 |  | ATOSU-58337-00011 | P0153 | CC_eb434 | 2229
PHSEND | 2013-04-25 |  | ATOSU-58337-00011 | P0153 | CC_ea578 | 101
PHSEND | 2013-04-25 |  | ATOSU-58337-00011 | P0153 | CC_eb578 | 38502
PHSEND | 2013-04-25 |  | ATOSU-58337-00011 | P0153 | CC_ind_slp | 1
PHSEND | 2013-04-25 |  | ATOSU-58337-00011 | P0153 | CC_ind_off | 0
PHSEND | 2013-04-25 |  | ATOSU-58337-00011 | P0153 | CC_psal | 35 | Constant
PHSEND | 2013-12-11 |  | ATOSU-58337-00012 | P0154 | CC_ea434 | 17533 |  | From CE04OSPS-00002
PHSEND | 2013-12-11 |  | ATOSU-58337-00012 | P0154 | CC_eb434 | 2229
PHSEND | 2013-12-11 |  | ATOSU-58337-00012 | P0154 | CC_ea578 | 101
PHSEND | 2013-12-11 |  | ATOSU-58337-00012 | P0154 | CC_eb578 | 38502
PHSEND | 2013-12-11 |  | ATOSU-58337-00012 | P0154 | CC_ind_slp | 1
PHSEND | 2013-12-11 |  | ATOSU-58337-00012 | P0154 | CC_ind_off | 0
PHSEND | 2013-12-11 |  | ATOSU-58337-00012 | P0154 | CC_psal | 35 | Constant
PHSEND | 2015-06-01 |  | A01345 | P0155 | CC_ea434 | 17533 |  | From CP04OSSM-00002
PHSEND | 2015-06-01 |  | A01345 | P0155 | CC_ea578 | 101
PHSEND | 2015-06-01 |  | A01345 | P0155 | CC_eb434 | 2229
PHSEND | 2015-06-01 |  | A01345 | P0155 | CC_eb578 | 38502
PHSEND | 2015-06-01 |  | A01345 | P0155 | CC_ind_off | 0
PHSEND | 2015-06-01 |  | A01345 | P0155 | CC_ind_slp | 1
PHSEND | 2015-06-01 |  | A01345 | P0155 | CC_psal | 35 | Constant
PHSEND | 2015-06-01 |  | A01346 | P0156 | CC_ea434 | 17533 |  | From CP04OSSM-00002
PHSEND | 2015-06-01 |  | A01346 | P0156 | CC_ea578 | 101
PHSEND | 2015-06-01 |  | A01346 | P0156 | CC_eb434 | 2229
PHSEND | 2015-06-01 |  | A01346 | P0156 | CC_eb578 | 38502
PHSEND | 2015-06-01 |  | A01346 | P0156 | CC_ind_off | 0
PHSEND | 2015-06-01 |  | A01346 | P0156 | CC_ind_slp | 1
PHSEND | 2015-06-01 |  | A01346 | P0156 | CC_psal | 35 | Constant
PHSEND | 2016-04-01 |  | A01346 | P0156 | CC_ea434 | 17533 |  | From CP01CNSM-00005
PHSEND | 2016-04-01 |  | A01346 | P0156 | CC_ea578 | 101
PHSEND | 2016-04-01 |  | A01346 | P0156 | CC_eb434 | 2229
PHSEND | 2016-04-01 |  | A01346 | P0156 | CC_eb578 | 38502
PHSEND | 2016-04-01 |  | A01346 | P0156 | CC_ind_off | 1
PHSEND | 2016-04-01 |  | A01346 | P0156 | CC_ind_slp | 0
PHSEND | 2016-04-01 |  | A01346 | P0156 | CC_psal | 35 | Constant
PHSEND | 2015-06-01 |  | A01342 | P0157 | CC_ea434 | 17533 |  | From CP03ISSM-00003
PHSEND | 2015-06-01 |  | A01342 | P0157 | CC_ea578 | 101
PHSEND | 2015-06-01 |  | A01342 | P0157 | CC_eb434 | 2229
PHSEND | 2015-06-01 |  | A01342 | P0157 | CC_eb578 | 38502
PHSEND | 2015-06-01 |  | A01342 | P0157 | CC_ind_off | 1
PHSEND | 2015-06-01 |  | A01342 | P0157 | CC_ind_slp | 0
PHSEND | 2015-06-01 |  | A01342 | P0157 | CC_psal | 35 | Constant
PHSEND | 2015-06-01 |  | A01343 | P0158 | CC_ea434 | 17533 |  | From CP4OSSM-00003
PHSEND | 2015-06-01 |  | A01343 | P0158 | CC_ea578 | 101
PHSEND | 2015-06-01 |  | A01343 | P0158 | CC_eb434 | 2229
PHSEND | 2015-06-01 |  | A01343 | P0158 | CC_eb578 | 38502
PHSEND | 2015-06-01 |  | A01343 | P0158 | CC_ind_off | 1
PHSEND | 2015-06-01 |  | A01343 | P0158 | CC_ind_slp | 0
PHSEND | 2015-06-01 |  | A01343 | P0158 | CC_psal | 35 | Constant
PHSEND | 2015-05-14 |  | A01426 | P0159 | CC_ea434 | 17533 |  | From CP4OSSM-00003
PHSEND | 2015-05-14 |  | A01426 | P0159 | CC_ea578 | 101
PHSEND | 2015-05-14 |  | A01426 | P0159 | CC_eb434 | 2229
PHSEND | 2015-05-14 |  | A01426 | P0159 | CC_eb578 | 38502
PHSEND | 2015-05-14 |  | A01426 | P0159 | CC_ind_off | 1
PHSEND | 2015-05-14 |  | A01426 | P0159 | CC_ind_slp | 0
PHSEND | 2015-05-14 |  | A01426 | P0159 | CC_psal | 35 | Constant
PHSEND | 2015-05-14 |  | A01427 | P0160 | CC_ea434 | 17533 |  | From CE06ISSM_00003
PHSEND | 2015-05-14 |  | A01427 | P0160 | CC_ea578 | 101
PHSEND | 2015-05-14 |  | A01427 | P0160 | CC_eb434 | 2229
PHSEND | 2015-05-14 |  | A01427 | P0160 | CC_eb578 | 38502
PHSEND | 2015-05-14 |  | A01427 | P0160 | CC_ind_off | 0
PHSEND | 2015-05-14 |  | A01427 | P0160 | CC_ind_slp | 1
PHSEND | 2015-05-14 |  | A01427 | P0160 | CC_psal | 35
PHSENA | 2013-02-07 |  | ATAPL-58337-00009 | P0161 | CC_ea434 | 17533 |  | From RS03AXPS
PHSENA | 2013-02-07 |  | ATAPL-58337-00009 | P0161 | CC_eb434 | 2229
PHSENA | 2013-02-07 |  | ATAPL-58337-00009 | P0161 | CC_ea578 | 101
PHSENA | 2013-02-07 |  | ATAPL-58337-00009 | P0161 | CC_eb578 | 38502
PHSENA | 2013-02-07 |  | ATAPL-58337-00009 | P0161 | CC_ind_slp | 1
PHSENA | 2013-02-07 |  | ATAPL-58337-00009 | P0161 | CC_ind_off | 0
PHSENA | 2013-02-07 |  | ATAPL-58337-00009 | P0161 | CC_psal | 35
PHSENA | 2013-02-07 |  | ATAPL-58337-00010 | P0162 | CC_ea434 | 17533 | From vendor cal sheet | From RS01SBPS
PHSENA | 2013-02-07 |  | ATAPL-58337-00010 | P0162 | CC_eb434 | 2229 | From vendor cal sheet
PHSENA | 2013-02-07 |  | ATAPL-58337-00010 | P0162 | CC_ea578 | 101 | From vendor cal sheet
PHSENA | 2013-02-07 |  | ATAPL-58337-00010 | P0162 | CC_eb578 | 38502 | From vendor cal sheet
PHSENA | 2013-02-07 |  | ATAPL-58337-00010 | P0162 | CC_ind_slp | 1 | Apply to all CGSN and RSN pH sensors; provided by the vendor as part of their Matlab code; DPS author wanted set as input to the code.
PHSENA | 2013-02-07 |  | ATAPL-58337-00010 | P0162 | CC_ind_off | 0 | Apply to all CGSN and RSN pH sensors; provided by the vendor as part of their Matlab code; DPS author wanted set as input to the code.
PHSENA | 2013-02-07 |  | ATAPL-58337-00010 | P0162 | CC_psal | 35 | Default per <ph_calc_phwater(ref, light, therm, ea434, eb434, ea578, eb578, ind_slp, ind_off, psal=35.0)>
PHSEND | 2015-08-31 |  | A01746 | P0167 | CC_ea434 | 17533 |  | From CE09OSSM-00002
PHSEND | 2015-08-31 |  | A01746 | P0167 | CC_ea578 | 101
PHSEND | 2015-08-31 |  | A01746 | P0167 | CC_eb434 | 2229
PHSEND | 2015-08-31 |  | A01746 | P0167 | CC_eb578 | 38502
PHSEND | 2015-08-31 |  | A01746 | P0167 | CC_ind_off | 0
PHSEND | 2015-08-31 |  | A01746 | P0167 | CC_ind_slp | 1
PHSEND | 2015-08-31 |  | A01746 | P0167 | CC_psal | 35
PHSEND | 2015-08-31 |  | OL000433 | P0168 | CC_ea434 | 17372 |  | From CE07SHSM-00003
PHSEND | 2015-08-31 |  | OL000433 | P0168 | CC_ea578 | 94.1
PHSEND | 2015-08-31 |  | OL000433 | P0168 | CC_eb434 | 2284
PHSEND | 2015-08-31 |  | OL000433 | P0168 | CC_eb578 | 38676
PHSEND | 2015-08-31 |  | OL000433 | P0168 | CC_ind_off | 0
PHSEND | 2015-08-31 |  | OL000433 | P0168 | CC_ind_slp | 1
PHSEND | 2015-08-31 |  | OL000433 | P0168 | CC_psal | 35
[9 empty] | PHSENF | 2015-08-10 | N00524 | P0170 | CC_ea434
[9 empty] | PHSENF | 2015-08-10 | N00524 | P0170 | CC_eb434
[9 empty] | PHSENF | 2015-08-10 | N00524 | P0170 | CC_ea578
[9 empty] | PHSENF | 2015-08-10 | N00524 | P0170 | CC_eb578
[9 empty] | PHSENF | 2015-08-10 | N00524 | P0170 | CC_ind_slp
[9 empty] | PHSENF | 2015-08-10 | N00524 | P0170 | CC_ind_off
[9 empty] | PHSENF | 2015-08-10 | N00524 | P0170 | CC_psal
PHSENF | 2015-08-10 |  | A01699 | P0171 | CC_ea434 | 17533 |  | From GS01SUMO-00002
PHSENF | 2015-08-10 |  | A01699 | P0171 | CC_eb434 | 2229
PHSENF | 2015-08-10 |  | A01699 | P0171 | CC_ea578 | 101
PHSENF | 2015-08-10 |  | A01699 | P0171 | CC_eb578 | 38502
PHSENF | 2015-08-10 |  | A01699 | P0171 | CC_ind_slp | 1
PHSENF | 2015-08-10 |  | A01699 | P0171 | CC_ind_off | 0
PHSENF | 2015-08-10 |  | A01699 | P0171 | CC_psal | 35 | Constant
PHSENF | 2015-06-01 |  | A01700 | P0172 | CC_ea434 | 17533 |  | From GS01SUMO-00002
PHSENF | 2015-06-01 |  | A01700 | P0172 | CC_eb434 | 2229
PHSENF | 2015-06-01 |  | A01700 | P0172 | CC_ea578 | 101
PHSENF | 2015-06-01 |  | A01700 | P0172 | CC_eb578 | 38502
PHSENF | 2015-06-01 |  | A01700 | P0172 | CC_ind_slp | 1
PHSENF | 2015-06-01 |  | A01700 | P0172 | CC_ind_off | 0
PHSENF | 2015-06-01 |  | A01700 | P0172 | CC_psal | 35 | Constant
PHSENF | 2015-06-01 |  | A01641 | P0173 | CC_ea434 | 17533 |  | From GI03FLMB-00002
PHSENF | 2015-06-01 |  | A01641 | P0173 | CC_eb434 | 2229
PHSENF | 2015-06-01 |  | A01641 | P0173 | CC_ea578 | 101
PHSENF | 2015-06-01 |  | A01641 | P0173 | CC_eb578 | 38502
PHSENF | 2015-06-01 |  | A01641 | P0173 | CC_ind_slp | 1
PHSENF | 2015-06-01 |  | A01641 | P0173 | CC_ind_off | 0
PHSENF | 2015-06-01 |  | A01641 | P0173 | CC_psal | 35 | Constant
PHSENF | 2015-06-01 |  | A01642 | P0174 | CC_ea434 | 17533 |  | From GI03FLMA-00002
PHSENF | 2015-06-01 |  | A01642 | P0174 | CC_eb434 | 2229
PHSENF | 2015-06-01 |  | A01642 | P0174 | CC_ea578 | 101
PHSENF | 2015-06-01 |  | A01642 | P0174 | CC_eb578 | 38502
PHSENF | 2015-06-01 |  | A01642 | P0174 | CC_ind_slp | 1
PHSENF | 2015-06-01 |  | A01642 | P0174 | CC_ind_off | 0
PHSENF | 2015-06-01 |  | A01642 | P0174 | CC_psal | 35 | Constant
PHSENE | 2015-06-01 |  | A01736 | P0176 | CC_ea434 | 17533 |  | From GS01SUMO-00002
PHSENE | 2015-06-01 |  | A01736 | P0176 | CC_eb434 | 2229
PHSENE | 2015-06-01 |  | A01736 | P0176 | CC_ea578 | 101
PHSENE | 2015-06-01 |  | A01736 | P0176 | CC_eb578 | 38502
PHSENE | 2015-06-01 |  | A01736 | P0176 | CC_ind_slp | 1
PHSENE | 2015-06-01 |  | A01736 | P0176 | CC_ind_off | 0
PHSENE | 2015-06-01 |  | A01736 | P0176 | CC_psal | 35 | Constant
PHSENE | 2015-06-01 |  | A01737 | P0177 | CC_ea434 | 17533 |  | From GS01SUMO-00002
PHSENE | 2015-06-01 |  | A01737 | P0177 | CC_eb434 | 2229
PHSENE | 2015-06-01 |  | A01737 | P0177 | CC_ea578 | 101
PHSENE | 2015-06-01 |  | A01737 | P0177 | CC_eb578 | 38502
PHSENE | 2015-06-01 |  | A01737 | P0177 | CC_ind_slp | 1
PHSENE | 2015-06-01 |  | A01737 | P0177 | CC_ind_off | 0
PHSENE | 2015-06-01 |  | A01737 | P0177 | CC_psal | 35 | Constant
PHSEND | 2015-06-01 |  | OL000322 | P00185 | CC_ea434 | 17372 |  | From CE02SHSM-00003
PHSEND | 2015-06-01 |  | OL000322 | P00185 | CC_ea578 | 94.1
PHSEND | 2015-06-01 |  | OL000322 | P00185 | CC_eb434 | 2284
PHSEND | 2015-06-01 |  | OL000322 | P00185 | CC_eb578 | 38676
PHSEND | 2015-06-01 |  | OL000322 | P00185 | CC_ind_off | 0
PHSEND | 2015-06-01 |  | OL000322 | P00185 | CC_ind_slp | 1
PHSEND | 2015-06-01 |  | OL000322 | P00185 | CC_psal | 35 | Constant
PHSEND | 2016-01-01 |  | OL000435 | P00187 | CC_ea434 | 17372 |  | From CE07SHSM-00003
PHSEND | 2016-01-01 |  | OL000435 | P00187 | CC_ea578 | 94.1
PHSEND | 2016-01-01 |  | OL000435 | P00187 | CC_eb434 | 2284
PHSEND | 2016-01-01 |  | OL000435 | P00187 | CC_eb578 | 38676
PHSEND | 2016-01-01 |  | OL000435 | P00187 | CC_ind_off | 0
PHSEND | 2016-01-01 |  | OL000435 | P00187 | CC_ind_slp | 1
PHSEND | 2016-01-01 |  | OL000435 | P00187 | CC_psal | 35 | Constant
Note on G374: smells like a roundoff issue -Dan Mergens
Note on G630: potential roundoff error? -Dan Mergens